## Sheet2 (hidden)
MAIN SWITCH TEST
120 | Measured Voltage |  |  | YES
132 | High VAC |  |  | NO
88 | Low VAC
Yes | Results Within Allowable Limits?
BASE CONNECTOR TO H1 CONTINUITY TEST
YES | 120 VAC Input Voltage Available from Cord Assembly?
YES | 120 VAC Input Voltage Available from Base Connector to HN1 Connector?
BATTERY SYSTEM TEST
Battery #1 Voltage Test
15 | Measured Voltage
14 | High VDC
11 | Low VDC
No | Results Within Allowable Tolerances?
Battery #2 Voltage Test
12 | Measured Voltage
14 | High VDC
11 | Low VDC
Yes | Results Within Allowable Tolerances?
CAPACITOR / RECTIFIER / AC120V TRANSFORMER TESTS
HN2 Connector Test (Pins 1 [red] & 3 [brown])
23 | Measured Voltage
23.5 | High Voltage DC
22.5 | Low Voltage DC
Yes | Results Within Allowable Tolerances?
HN2 Connector Test (Pins 3 [orange] & 5 [black (-)])
23 | Measured Voltage
27.6 | High Voltage DC
26.4 | Low Voltage DC
No | Results Within Allowable Tolerances?
CHARGING INDICATOR TEST
Test Output from Relay Box at CN3 Pin # 11
23 | Measured Voltage
5.5 | High Voltage DC
4.5 | Low Voltage DC
No | Results Within Allowable Tolerances?
PENDANT CONTROL CORD TEST
Test for continuity across pins from the base end to the pendant control end.
| BASE END PIN # | PENDANT END PIN # | Continuity Obtained?
| 1- Black | 1- Black
| 2- Brown | 2- Brown | YES
| 3- Red | 3- Red | YES
| 4- Green | 4- Green
| 5- Blue | 5- Blue | YES
| 6- White | 6- White | YES
TABLE BASE PENDANT CONTROL CONNECTOR TO RELAY BOX (HN8) TEST
Test for continuity across pins from the base end to the pendant control end.
| BASE END PIN # | CN7 PIN # | Continuity Obtained?
| 1- Black | 1- Black
| 2- Brown | 2- Brown
| 3- Red | 3- Red
| 4- Green | 4- Green
| 5- Blue | 5- Blue | YES
| 6- White | 6- White | YES
| 7- UNUSED | 7- UNUSED

## Electrical (Battery)
Table Model Number :
Table Serial Number:
Hospital:
Technician:
Date:
Describe Issue:
MAIN POWER TEST (VAC) |  |  | Comments
0 | What is the voltage at the wall outlet?
132 | High VAC |  |  | YES
88 | Low VAC |  |  | NO
No | Results Within Allowable Limits?
0 | What is the voltage coming through the power cord?
132 | High VAC
88 | Low VAC
No | Results Within Allowable Limits?
0 | What is the output voltage of the main power switch at the HN1 connector?
132 | High VAC
88 | Low VAC
No | Results Within Allowable Limits?
0
0 | What is the measured voltage of battery #1? |  | Out of Specification
13 | High VDC
11 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the measured voltage of battery #2? |  | Out of Specification
13 | High VDC
11 | Low VDC
No | Results Within Allowable Tolerances?
0
0 | What is the output voltage of the transformer at the HN2 connector?
23 | High VAC
21 | Low VAC
No | Results Within Allowable Tolerances?
0 | What is the voltage load of the cap/rec unit at the HN2 connector? |  | Out of Specification
23.5 | High VAC
22.5 | Low VAC
No | Results Within Allowable Tolerances?
0
0 | What is the voltage output of the cap/rec unit to the relay box at the HN3 connector?
31 | High VDC
29 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the voltage output of the cap/rec unit to the batteries at the HN3 connector?
27.6 | High VDC
26.4 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the voltage output from the relay box to the charge indicator at the CN3 connector?
5.5 | High VDC
4.5 | Low VDC
No | Results Within Allowable Tolerances?
0
0 | What is the input voltage to the relay box in battery mode at the CN4 connector?
28 | High VDC
24 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the input voltage to the relay box in AC120V mode at the CN4 connector?
28 | High VDC
24 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the pump/motor at the HN4 connector?
28 | High VDC
24 | Low VDC
No | Results Within Allowable Tolerances?
0
0 | What is the output voltage to the TABLE UP solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the TABLE DOWN solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the TREND solenoid at the CN1A connector?
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the REV TREND solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the TILT RIGHT solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the TILT LEFT solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the REFLEX solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the FLEX solenoid at the CN1A connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the BACK UP solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the BACK DOWN solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the SLIDE TO FOOT solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the SLIDE TO HEAD solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the KIDNEY DOWN solenoid at the CN1B connector?
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the KIDNEY UP solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the LEG UP solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the LEG DOWN solenoid at the CN1B connector?
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the BRAKE LOCK solenoid at the CN1B connector? |  | Out of Specification
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0 | What is the output voltage to the BRAKE UNLOCK solenoid at the CN1B connector?
24.5 | High VDC
23.5 | Low VDC
No | Results Within Allowable Tolerances?
0
0 | What is the resistance value of the TABLE UP solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the TABLE DOWN solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the TREND solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the REV TREND solenoid at the CN1A connector? |  | Out of Specification
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the TILT RIGHT solenoid at the CN1A connector? |  | Out of Specification
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the TILT LEFT solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the REFLEX solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the FLEX solenoid at the CN1A connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the BACK UP solenoid at the CN1B connector? |  | Out of Specification
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the BACK DOWN solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the SLIDE TO FOOT solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the SLIDE TO HEAD solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the KIDNEY DOWN solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the KIDNEY UP solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the LEG UP solenoid at the CN1B connector? |  | Out of Specification
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the LEG DOWN solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the BRAKE LOCK solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the BRAKE UNLOCK solenoid at the CN1B connector?
20 | High Ω
16 | Low Ω
No | Results Within Allowable Tolerances?
0
0 | What is the input voltage to the pump/motor at HN4?
28 | High Ω
24 | Low Ω
No | Results Within Allowable Tolerances?
0 | What is the resistance value of the pump/motor at HN4 (Ω)? |  | Out of Specification
2 | High Ω
1 | Low Ω
No | Results Within Allowable Tolerances?
0
0
0
0 | Pendant End Pin # | Base End Pin #
| 1- Black | 1- Black
| 2- Brown | 2- Brown
| 3- Red | 3- Red
| 4- Green | 4- Green
| 5- Blue | 5- Blue
| 6- White | 6- White
TABLE PENDANT CONTROL CORD (HN8) to RELAY BOX TEST (CN7)
Continutity Obtained? | Pendant Connection End Pin # | CN7 Pin #
| 1- Black | 1- Black
| 2- Brown | 2- Brown
| 3- Red | 3- Red
| 4- Green | 4- Green
| 5- Blue | 5- Blue
| 6- White | 6- White
| 7- UNUSED | 7- UNUSED
TABLE FOOT CONTROL CORD (HN9) to RELAY BOX TEST (CN8)
Continutity Obtained? | Foot Control Connection End Pin # | CN8 Pin #
| 1- Black | 1- Black
| 2- Brown | 2- Brown
| 3- Red | 3- Red
| 4- Green | 4- Green
| 5- Blue | 5- Blue
| 6- White | 6- White
| 7- UNUSED | 7- UNUSED
AUXILIARY TOGGLE SWITCH to RELAY BOX TEST (CN2)
Continutity Obtained? | Function | CN2 Pin Numbers
| Reverse Trend | 16/4 | Out of Specification
| Trend | 16/3 | Out of Specification
| Tilt Right | 16/5 | Out of Specification
| Tilt Left | 16/6 | Out of Specification
| Back Up | 16/7 | Out of Specification
| Back Down | 16/8 | Out of Specification
| Table Up | 16/1 | Out of Specification
| Table Down | 16/2 | Out of Specification
| Leg Up | 16/13 | Out of Specification
| Leg Down | 16/14 | Out of Specification
| Brake Unlock | 16/12 | Out of Specification
| Kidney Down | 16/11 | Out of Specification
MAIN WIRE HARNESS to RELAY BOX TEST (CN4)
Continutity Obtained? | CN4 Pin Number | Connector
| 1-Red | Battery
| 2-Blue | Battery
| 3-Red | HN4
| 4-Black | HN4
| 5-Purple | HN3
| 6-Black | HN3
| 7-White | HN4
| 8-White | HN4
| 9-Yellow | HN3
| 10-Gray | HN3
| 11-Orange | HN3
| 12-NOT USED | ---

## Block Diagram
SYMPTOMS
All Control Functions are Affected | One Control Function is Affected | All or One Control Function is Affected
Motor / Pump Assy. | Solenoid | Pendant Control
Main Switch Circuit |  | Relay Box Assy
Wiring |  | Auxillary Switches
Pendant Control |  | Batteries
Batteries |  | Return-to-Level Switches
Capacitor Rectifier

## Results (hidden)
MAIN POWER
No | Is the Wall Outlet Good?
No | Is the Power Cord Good?
No | Is the Main Power Switch Good?
BATTERIES
No | Is Battery #1 Good?
No | Is Battery #2 Good?
CHARGING SYSTEM
No | Is the Transformer Good?
No | Does the Cap/Rec Unit Provide the Proper Load?
No | Does the Cap/Rec Unit Provide the Voltage to the Relay Box?
No | Does the Cap/Rec Unit Provide the Voltage to the Batteries?
No | Is the Voltage correct from the Relay Box to the Charge Indicator?
RELAY BOX
No | Is the input voltage correct in battery mode?
No | Is the input voltage correct in AC 120V mode?
No | Is the output voltage correct to the pump/motor?
No | Is the output voltage correct to the TABLE UP solenoid?
No | Is the output voltage correct to the TABLE DOWN solenoid?
No | Is the output voltage correct to the TREND solenoid?
No | Is the output voltage correct to the REV TREND solenoid?
No | Is the output voltage correct to the TILT RIGHT solenoid?
No | Is the output voltage correct to the TILT LEFT solenoid?
No | Is the output voltage correct to the REFLEX solenoid?
No | Is the output voltage correct to the FLEX solenoid?
No | Is the output voltage correct to the BACK UP solenoid?
No | Is the output voltage correct to the BACK DOWN solenoid?
No | Is the output voltage correct to the SLIDE TO FOOT solenoid?
No | Is the output voltage correct to the SLIDE TO HEAD solenoid?
No | Is the output voltage correct to the KIDNEY DOWN solenoid?
No | Is the output voltage correct to the KIDNEY UP solenoid?
No | Is the output voltage correct to the LEG UP solenoid?
No | Is the output voltage correct to the LEG DOWN solenoid?
No | Is the output voltage correct to the BRAKE LOCK solenoid?
No | Is the output voltage correct to the BRAKE UNLOCK solenoid?
SOLENOIDS
No | Is the resistance value correct to the TABLE UP solenoid?
No | Is the resistance value correct to the TABLE DOWN solenoid?
No | Is the resistance value correct to the TREND solenoid?
No | Is the resistance value correct to the REV TREND solenoid?
No | Is the resistance value correct to the TILT RIGHT solenoid?
No | Is the resistance value correct to the TILT LEFT solenoid?
No | Is the resistance value correct to the REFLEX solenoid?
No | Is the resistance value correct to the FLEX solenoid?
No | Is the resistance value correct to the BACK UP solenoid?
No | Is the resistance value correct to the BACK DOWN solenoid?
No | Is the resistance value correct to the SLIDE TO FOOT solenoid?
No | Is the resistance value correct to the SLIDE TO HEAD solenoid?
No | Is the resistance value correct to the KIDNEY DOWN solenoid?
No | Is the resistance value correct to the KIDNEY UP solenoid?
No | Is the resistance value correct to the LEG UP solenoid?
No | Is the resistance value correct to the LEG DOWN solenoid?
No | Is the resistance value correct to the BRAKE LOCK solenoid?
No | Is the resistance value correct to the BRAKE UNLOCK solenoid?
PUMP/MOTOR
No | Does the pump/motor have the correct input voltage?
No | Does the pump/motor have the correct resistance value?